## Slide 1: Bug in lcscope.dll
Marcel Schumacher, Petio Natzkin

## Slide 2: Bug Description
The current LeCroy scope driver with version 3.2.9.0 contains a bug that mostly but not only affect 32-bit systems
On 64-bit the bug could not be verified so far, but generally a more stable bahavior could be observed by the reporters
Patched DLLs are provided here fixing the bug, so please copy the DLLs into the IVI installation directory after driver installation and replace the existing onces
  lcscope.dll  C:\Program Files (x86)\IVI Foundation\IVI\Bin
  lcscope_64.dll  C:\Program Files\IVI Foundation\IVI\Bin
More information about the bug can be found on following slides

## Slide 3: Basics:
Oscilloscope can set resolution by providing a minimum and a maximum value
CFMT DEF9,WORD,BIN -> Data is sent as 16-bit signed integers but only the MSB is used. The configured minimum is represented by (0x8000) and the maximum by (0x7F00). Since LSB is unused for simplification we can say our values range from -128 (0x80) to 127 (0x7F)
CORD LO -> Data is sent in little endian format (<LSB><MSB>)
Source-code of lcscope.dll is provided by teledyne and can be installed along with the driver but we could not compile it ourselves for lack of Visual Studio
  (https://teledynelecroy.com/support/softwaredownload/download.aspx?did=7514)

## Slide 4: Bug description:
Issue:
  When calling FetchWaveform(…), the program hangs and waits forever
Appearance:
  Bug does not appear for the first call to FetchWaveform(…)
  Could only be triggered in the 32 bit version of the dll (i.e. must use 32 bit python); In theory, the 64 bit version contains the same bug but it did not appear for us because of a different memory layout
  Happens, when a value of 0x0A is transferred. This corresponds to , where
    E.g. setting Min to -24.5 and Max to 20, 0x0A represents -0.5±0.2 and thus the value should be read, when no voltage is applied

## Slide 5: Trace of correct call (captured with NI I/O Trace on x64)
Header says 1004 bytes are available and all are read with one call to viRead(…) (1021 = 1004 + sizeof(header) + Termchar)
Termchar was set to disabled and is not modified (255, 256)
Buffer contains 16 bit-values, can contain 0x0A00 (Little Endian) as a value and ends with a single 0x0A (line feed, default Termchar)

## Slide 6: Trace of error (1/2) (captured with NI I/O Trace on x32)
First call of FetchWaveform(…)
During first call, everything is read like before but after that, TERMCHAR is read, found to be FALSE (115), set to FALSE (116) and set to TRUE (117).
After that, the TERMCHAR is read as TRUE (118), set to FALSE (119) and set to TRUE again (120)
 Before the call, the TERMCHAR is set to disabled (115), while it is enabled after it (120)

## Slide 7: Trace of error (2/2) (captured with NI I/O Trace on x32)
Second call to FetchWaveform(…):
Only 26 bytes (10 data-bytes) are read, although 1002 data-bytes are promised (187) and code 0x3fff0005 (VI_SUCCESS_TERM_CHAR) is returned
We can see, the last read character is the first 0x0A, which is actually part of the data and not the termination character
This seems logical, since we can see that the termination character was enabled (188) (we saw on the last slide that it was erroneously enabled)
Termination character is then disabled (189) and the read is tried again (190), whereas no data are read. The same is tried twice again (194 and 198), whereas 198 waits indefinitely, since no more data are available

## Slide 8: Correct behaviour using VISA directly (captured with NI I/O Trace on x32)
The same communication between the device and the software was rebuilt, using pyvisa to communicate directly with the device (i.e. without the use of lcscope.dll)
Only 42 bytes (26 data-bytes) are read, although 1004 data-bytes are promised (189) and code 0x3fff0005 (VI_SUCCESS_TERM_CHAR) is returned
We can see, the last read character is the first 0x0A, which is actually part of the data and not the termination character
Termination character is then disabled (190) and the read is tried again (192), whereas the remaining 979 bytes of data are read.
After that, the termchar is toggled as before

## Slide 9: Two bugs
Termination character should not have been enabled in the first place!
The data was discarded after the termination character was read (in the pyvisa case we could see that the data could be read after the termchar was disabled, while in the lcscope.dll-case, the data was not there anymore)

## Slide 10: Why was termcharacter enabled?
Excerpt from NTL\NtlIo.inl (used for lcscope.dll):
inline HRESULT CInstrBuffer::ReadBinaryBlockData(BYTE* pBuf, long cbDesired, long* pcbActual, bool bFlush)
{
// (...)
HRESULT hr = S_OK;
// Disable the termchar if it's enabled
bool bTermCharEnabled;
hr = m_pIo->GetTerminationCharacterEnabled(&bTermCharEnabled);
if (SUCCEEDED(hr))
{
if (bTermCharEnabled)
hr = m_pIo->SetTerminationCharacterEnabled(false);
if (SUCCEEDED(hr))
{
// read data
//(...)
// Re-enable the termchar if previously enabled
if (bTermCharEnabled)
{
HRESULT hr2 = m_pIo->SetTerminationCharacterEnabled(true);
hr = FAILED(hr) ? hr : hr2;
}
}
}
return hr;
}
Must have been erroneously set to true since code below is executed

## Slide 11: Why did GetTerminationCharacterEnabled() return true, although the trace showed false?
inline HRESULT CVisaCSession::GetAttribute(long lAttribute, bool* pbValue)
{
//...
*pbValue = false;
HRESULT hr = S_OK;
ViStatus status = VI_SUCCESS;
ViBoolean value;
status = viGetAttribute(m_session, lAttribute, &value);
hr = TranslateStatus(status);
if (SUCCEEDED(hr)) {
*pbValue = value != VI_FALSE ? true : false;
}
return hr;
}
Excerpt from NTL\NtlVisaC.inl:
Typedef to unsigned short (16 bit) in external include; Set to 0 by viGetAttribute
Compiled to:
cmp dword ptr ss:[ebp-10], 0
32 bit comparison of 16 bit variable
Only lower 16 bit are modified by viGetAttribute(…)
If other 16 bits equal 0, this expression is evaluated correctly, otherwise it is evaluated to true

## Slide 12: Why did GetTerminationCharacterEnabled() return true, although the trace showed false?
// Custom I/O provider is needed just so we can override the behavior of GetAttribute when boolean parameters
// are involved. VICP has a bug which expects a ViInt32 output parameter for boolean attribute queries.
// ...
virtual HRESULT GetAttribute(long lAttribute, bool* pbValue) override
{
//...
*pbValue = false;
HRESULT hr = S_OK;
ViStatus status = VI_SUCCESS;
// Pass in a ViInt32, even though we should pass in a ViBoolean. This is a bug in VICP
// that was likely copied from the same bug in NI-VISA.
ViInt32 value;
status = viGetAttribute(m_session, lAttribute, &value);
hr = TranslateStatus(status);
if (SUCCEEDED(hr)) {
*pbValue = value != VI_FALSE ? true : false;
}
return hr;
}
Excerpt from LeCroyScope\LecroyScopeIoProvider.h
(overwriting corresponding funciton from NTL\NtlVisaC.inl):
Uninitialized 32-bit variable for a boolean
32 bit comparison compiled to:
cmp dword ptr ss:[ebp-10], 0
If higher 16 bits equal 0, this expression is evaluated correctly, otherwise it is always evaluated to true, independent from the lower 16 bits
For x32 lcscope.dll, the higher 16 bits were unequal to 0  always evaluated to true
For x64 lcscope.dll, the higher 16 bits were (randomly?) equal to 0  evaluated correctly
Only lower 16 bit of „value“-variable are modified by viGetAttribute(…)  Apparently the bug mentioned in the comments was fixed in VICP
 Higher 16 bit remain uninitialized and are random

## Slide 13: Solution (1/2)
Patch lcscope.dll, so that
  cmp dword ptr ss:[ebp-10], 0
  cmp byte ptr ss:[ebp-10], 0
  (i.e. modify byte at file-offset 0x26ef4 from 0x83 to 0x80)
Compatible to versions of VICP with and without the bug mentioned in the code-comments
Strictly speaking, this is not entirely correct either, since actually „word“ would match 16 bits but the machine code for „cmp word ptr“ is one byte larger than that of „cmp dword ptr“ so there is not enough space without shifting everything. However, since the value should only contain 0 (VI_FALSE) or 1 (VI_TRUE), a one byte comparison is sufficient

## Slide 14: Solution (2/2)
This solution fixes the first of the two bugs, where the termination character was erroneously set
By doing so, the second bug never appears, since the termination character is never enabled
Open Question:
Why does the second bug appear (i.e. why can‘t the data be read after the termchar was explicitly disabled)? We saw this did work when using visa directly without lcscope.dll
  Was not further investigated, since it could be circumvented